TAbMEP Assessment: POLARCAT O3 Measurements

- Introduction

Here we provide the assessment for the ozone (O3) measurements taken from five aircraft platforms during the summer 2008 POLARCAT field campaign [INSERT REFERENCE].  This assessment is based upon three wing-tip-to-wing-tip intercomparison flights conducted during the field campaign.  The remaining four flights in the POLARCAT field campaign are left out of calculations due to pending data submissions.  Recommendations provided here offer a systematic approach to unifying the POLARCAT O3 data for any integrated analysis.  These recommendations are based upon the instrument performance demonstrated during the POLARCAT measurement comparison exercises and are not to be extrapolated beyond this campaign.

- POLARCAT O3 Measurements

Five different O3 instruments were deployed on the five aircraft.  Table 1 summarizes these techniques and gives references for more information.

Table 1.  O3 measurements deployed on aircraft during POLARCAT

| Aircraft | Instrument | Reference |
| --- | --- | --- |
| NASA DC-8 | CLD |  |
| NASA P-3B |  |  |
| NOAA WP-3D | CLD |  |
| DLR FALCON |  |  |
| ATR-42 FALCON |  |  |

- Summary of Results

Table 2 summarizes the recommendations drawn from the intercomparisons.  The following sections describe the processes that led to the recommendations.  Table 2 recommends a bias correction (see section 4.1 for details) that can be applied to each data set to maximize the consistency between them.  The recommended 2σ uncertainty in Table 2 is the larger of either the twice uncertainty reported by the PI or the quadrature-sum of the recommended bias correction listed in Table 2 and twice the adjusted precision determined for each instrument (see Table 4).  Where there is no uncertainty reported by the PI, the recommended 2σ uncertainty is the quadrature-sum of the recommended bias correction and twice the adjusted precision.  When there are multiple intercomparisons available for the same instrument, the maximum adjusted precision value is used.

Table 2. Recommended POLARCAT O3 measurement treatment

| Aircraft | Instrument | Reported 1σ Uncertainty | Recommended Bias Correctiona | Recommended 2σ Uncertainty |
| --- | --- | --- | --- | --- |
| NASA DC-8 |  |  | -1.19 + 0.0803 O3DC8 | {(-1.19 + 0.0803 O3)2 + (.052)2}1/2 |
| NASA P-3B |  |  |  |  |
| NOAA WP-3D |  | ± .05 + 4% | -2.2 - 0.135 O3WP3D | {(-2.2 - 0.135 O3)2 + (0.0148)2}1/2 |
| DLR FALCON |  |  | 1.73 - 0.230 O3DLR | {(1.73 - 0.230 O3)2 + (.096)2}1/2 |
| ATR-42 FALCON |  |  | 3.00 - 0.273 O3ATR | {(3.00 - 0.273 O3)2 + (.106)2}1/2 |

## a The “true O3 mixing ratio” = measurement – recommended bias correction (as discussed in Section 4.1).

- Results and Discussion

  - Bias Analysis

## Figures 1 – 3 illustrate the need for quantifying the bias between instruments.  The difference between the simultaneous measurements reported by two instruments is plotted against the O3 mixing ratio reported by one of the instruments.  The apparent biases in Table 3 are calculated from orthogonal linear regression (ODR) analysis (shown in the correlation plots in Figs. A1–A3).  ODR is used to approximate the bias between the paired instruments’ dependence on the O3 mixing ratio.  Apparent bias is defined as the difference in a measurement on one aircraft platform referenced to the same measurement made on the DC-8 (i.e. WP-3D – DC8).  For convenience, the apparent bias is given in the form a + b*O3 DC8.  In this form, it is easier to propagate the apparent biases so the best estimate bias can be used to calculate the uncertainties summarized in Table 2.  It should be noted here that the intercept should not simply be interpreted as a measurement offset; instead it is used in conjunction with the slope to best describe the linear trend found in the data.

## The best estimate bias is defined as the difference between the instrument being analyzed and the true O3 mixing ratio as a function of the instrument being analyzed.  This can be calculated by subtracting the true O3 mixing ratio from the respective apparent bias equation from Table 3 and expressing the result in terms of the instrument being analyzed.  The average of the apparent biases for four instruments (1.19 ppbv - 0.0803 O3 DC8) is assumed to be the best estimate of the “true O3 mixing ratio” from the DC-8 O3 measurement.  In effect, this procedure assumes that the best estimate of the true O3 mixing ratio is the average of the five instruments, and the apparent bias correction is used in calculations to most closely approximate the true O3 mixing ratio for each instrument.

## It should be noted that the initial choice of the reference instrument is arbitrary, and has no impact on the final recommendations.  The given bias corrections were based upon the instrument performance demonstrated during the intercomparison periods.

Table 3. POLARCAT O3 bias estimates

| Aircraft | Instrument | Apparent Bias1 (a ppbv + b O3) | Best Estimate Bias (a ppbv + b O3) |
| --- | --- | --- | --- |
| NASA DC-8 |  | 0 | -1.19 + 0.0803 O3DC8 |
| NASA P-3B |  |  |  |
| NOAA WP-3D |  | - 0.904 - 0.0480 O3DC8 | -2.2 - 0.135 O3WP3D |
| DLR FALCON |  | 2.37 - 0.122 O3DC8 | 1.73 - 0.230 O3DLR |
| ATR-42 FALCON |  | 3.29 - 0.151 O3DC8 | 3.00 - 0.273 O3ATR |

1 DC-8 is taken as an arbitrary reference.  Apparent bias is expressed as a linear function of O3 on the DC-8.

## 4.2 Precision Analysis

## The instrument precision assessment is summarized in Table 4.  The Internal Estimate of Instrument Precision (IEIP) analysis procedures were applied for all continuous fast instruments.  The IEIP procedure is an effective method to estimate “short-term” precision, which accounts for signal variation during a short period of assumed constant O3 measurements.  Because this assumption is not always valid, the IEIP estimate tends to provide an upper limit of the instrument short-term precision.  Over longer time scales, however, some instruments are subject to lower precision (i.e. larger variability), which includes variability that arises from uncorrected changes in the zero level or sensitivity of the instrument.  These additional contributions to the variability are not likely reflected in the IEIP derived precision, but the intercomparison flights do provide a reasonable check on their influence.  This effect was examined through the comparisons of the “expected variability" and "observed variability" given in Table 4.  The expected variability is the quadrature-sum of the corresponding IEIP precisions.  The observed variability is the standard deviation derived from the three intercomparisons shown in Figs. 4-6, denoting the relative difference between the paired instruments.  Each standard deviation is expected to be equal to the quadrature-sum of the separate IEIP precisions of the two intercompared instruments.  In two cases the observed variability is larger than the expected variability, which indicates that the IEIP derived (short-term) precision needs to be adjusted to reflect the longer term fluctuations.  Table 4 contains estimates of this “adjusted” precision obtained by proportionally scaling the IEIP estimates so that the expected variability values would equal to that of the observed variability.  For the case that the observed variability is smaller, the adjusted precision (last column in Table 4) is set equal to the IEIP precision.  Because IEIP precision could not be calculated for ATR-42 Falcon data, the 7/14 DLR Falcon IEIP precision value was taken as the DLR Falcon adjusted precision value and used with the observed variability to calculate the adjusted precision for the ATR-42 Falcon flight.  Based on the results presented in Table 4, the worst "adjusted precision" (or the largest value) is taken as a conservative precision estimate for each POLARCAT O3 instrument and is used for the derivation of the recommended 2σ uncertainty in the last column of Table 2.

## Table 3 shows that the measurement bias is a function of O3 mixing ratio.  Thus, the bias may have a significant impact on the observed variability.  To minimize the effect of bias, we make corrections for bias before computing the observed variability, but only when this reduces the variability.  For instance, the observed variability in the case of DLR/ATR-42 flight on 7/14 was estimated at 18.0% without correction.  This value was reduced to 5.39% when bias correction was applied.  The observed variability values given in Table 4 are computed after the bias correction. The final analysis results are shown in Table 2.  Over 90% of the data falls within the combined recommended uncertainties for each intercomparison, which is consistent with the TAbMEP guideline for unified data sets.

Table 4. POLARCAT O3 precision (1σ) comparisons

| Flight | Platform | IEIP Precision | Expected Variability | Observed Variability | Adjusted Precision |
| --- | --- | --- | --- | --- | --- |
| 4/8 | DC-8 |  |  |  |  |
|  | P-3B |  |  |  |  |
| 4/12 | DC-8 | 0.5% | 0.61% | 1.30% | 1.1% |
|  | WP-3D | 0.4% |  |  | 0.8% |
| 4/19 | DC-8 |  |  |  |  |
|  | P-3B |  |  |  |  |
| 7/9 | DC-8 | 1.7% | 3.54% | 5.44% | 2.6% |
|  | DLR Falcon | 3.1% |  |  | 4.8% |
| 7/10 | DC-8 |  |  |  |  |
|  | P-3B |  |  |  |  |
| 4/15 | P-3B |  |  |  |  |
|  | WP-3D |  |  |  |  |
| 7/14 | DLR Falcon | 1.1% |  | 5.39% | 1.1% |
|  | ATR-42 Falcon |  |  |  | 5.3% |

## Appendix A

Figures A1 through A3 show the time series of the O3 measurements and aircraft altitudes for each intercomparison flight as well as the correlations between the two O3 measurements.

## References

[INSERT REFERENCES]

Figures

Figure 1: Difference between O3 measurements for the DC-8/WP-3D intercomparison flight as a function of DC-8 O3.

Figure 2: Difference between O3 measurements for the DC-8/DLR Falcon intercomparison flight as a function of DC-8 O3.

Figure 3: Difference between O3 measurements for the DLR/ATR-42 Falcon intercomparison flight as a function of DLR O3.

Figure 4: Relative difference between O3 measurements from the DC-8/WP-3D intercomparison flight as a function of DC-8 O3.

Figure 5: Relative difference between O3 measurements from the DC-8/DLR Falcon intercomparison flight as a function of DC-8 O3.

Figure 6: Relative difference between O3 measurements from the DLR/ATR-42 Falcon intercomparison flight as a function of DLR O3.  Corrections were made to the ATR-42 Falcon data to account for bias in the correlation with DLR.

Appendix

Figure A1: (left panels) Time series of O3 measurements and aircraft altitudes from two aircraft on the intercomparison flight between NASA DC-8 and NOAA WP-3D.

(right panels) Correlations between the O3 measurements on the two aircraft.

Figure A2: (left panels) Time series of O3 measurements and aircraft altitudes from two aircraft on the intercomparison flight between NASA DC-8 and DLR Falcon.

(right panels) Correlations between the O3 measurements on the two aircraft.

Figure A3: (left panel) Time series of O3 measurements and aircraft altitudes from two aircraft on the intercomparison flight between DLR Falcon and ATR-42 Falcon.

(right panel) Correlations between the O3 measurements on the two aircraft.